## Slide 1
The BigBOSS Fiber View Camera
1
Progress Report
C Baltay
July 18, 2013

## Slide 2: The BigBOSS Fiber View Camera
A collaborative effort between LBNL and Yale University
  C. Baltay, W. Emmet, and N. Padmanabhan, B. Horowitz, Terry Girard*, Dana Cassetti*, Yale
  R. Besuner, M. Lampton, and M. Sholl, LBNL
  *members of Prof. Bill Van Altena’s precision astrometry group at Yale

## Slide 3: The Function of the Fiber View Camera
5000 fibers on the fiber plane will be positioned to their desired positions by a computer controlled mechanism. Expected precision of fiber positioningis expected to be 50 to 100 microns. Would like 5 micron accuracy.
The fiber View Camera is designed to take an image of the fiber plane to measure the position of each fiber after the mechanism is done positioning them.
The positions of misplaced fibers can then be corrected based on the information from the Fiber View Camera.

## Slide 4: Fiber View Camera Draft Requirements
Feature Requirement Comment
Absolute Position Meas. 5 μ Needs Test
Nearest Neighbor Dist. 3 mm Easy
Exposure Time < 2 sec Easy
Readout Time < 2 sec Easy

## Slide 5: Some Design considerations
Input dimensions
  Fiberplane diameter 1000 mm
  CCD size ~40 mm
  Fiberplane to camera L = 8900 mm
Demagnification desired ~ 25
Demag ~ f/L
Camera focal length f ~ 400 mm
At a demagnification of 25, the 5 micron precision requirement on fiber plane translates to 0.2 microns on CCD, or 1/30 of a 6 micron pixel (33 millipixels)

## Slide 6: Fiber View Camera R&D Plan
Goal of the R&D Program is to demonstrate that the required 5 micron precision on the fiber position is achievable. To do this we plan to:
  Design and build a prototype camera
  Design and build a test setup at Yale with a mock up of a portion of the fiber plane with fibers with precisely measured positions
  Carry out the test by photographing the mock fiber plane and analyzing the results

## Slide 7: The Mayall 4 m Telescope Modified Optics for BigBOSS
Fiber Plane
Fiber View Camera
Corrector Optics
Primary Mirror
8.9 m
2.5 m

## Slide 8: Mounting of the Fiber View Camera
This is the current baseline location for the FVC

## Slide 9: Illuminating the fibers
Fiber Plane
Spectrometer Enclosure
LED
Fibers
Fiber View Camera

## Slide 10: Lab Calibration/Test Setup
LED
Fibers
Plate with precision
Measured holes
Camera
Insert 120 micron fibers into an invar plate. We have at Yale an Avant
Optical measuring machine with ~micron precision and an area
Coverage of 18” by 24” to measure positions of these fibers.

## Slide 11: Choice of Prototype Camera Components
Best choices are:
  Kodak KAF 50100 CCD
  Canon EF 400 mm f/5.6 lens
  FLI Proline PL50100 Camera Electronics
Placed Orders March 31, 2012
Cost of the prototype camera:
  FLI camera including the CCD $ 21,045
  Canon 400 mm lens 1,249
  Canon Lens Adapter, focus control 981
  Total for prototype camera $ 23,275

## Slide 15: CCD, electronics, and lens assembled Completed May 1, 2012

## Slide 16: Canon Lens Adapter Focus and Aperture Computer controlled

## Slide 17: Image of Air Force Test Pattern

## Slide 18: Fiber Plane Simulation
Designed and Built a Fiber Plane Prototype
18” x 24” invar plate 1/8” thick
Fibers from Berkeley
  120 micron dia fibers
  2 bundles of 32 fibers each, 64 fibers total
18 fiducials, 46 randomly placed fibers
Illuminated by a 10mwatt LED

## Slide 19: Fiber Plane Prototype
Fiber Plane with 61
Fiducial Fibers
18” x 24” prototype
with 18 of the fiducials
X=0, y=0

## Slide 20: Fiducial Fibers
FVC Focal Plate Fiducials in mm
Numbe r theta x y
1 0 0 0.000 0.000
2 120 180 0.000 -120.000
3 120 120 103.923 -60.000
4 120 60 103.923 60.000
5 120 0 0.000 120.000
6 240 120 207.846 -120.000
7 240 90 240.000 0.000
8 240 60 207.846 120.000
9 240 30 120.000 207.846
10 240 0 0.000 240.000
11 360 100 354.531 -62.513
12 360 80 354.531 62.514
13 360 60 311.769 180.000
14 360 40 231.403 275.776
15 480 105 463.644 -124.233
16 480 90 480.000 0.000
17 480 75 463.644 124.233
18 480 60 415.692 240.000

## Slide 21: Fiber Plane Prototype 120 micron dia fibers from Berkeley

## Slide 22: Fiber Plane Prototype Using two fiber bundles, 32 fibers each

## Slide 23: Fiber Plane Prototype

## Slide 24: Fiber Plane Prototype 64 fibers total

## Slide 25: Fiber Plane Prototype 120 micron fibers in an invar plate

## Slide 26: Fiber Plane Prototype 18 fiducial fibers, 46 randomly placed fibers

## Slide 27: OGP High Precision Measuring Machine
OGP (Optical Gaugeing Products, Rochester,NY) Model Avant ZIP 400/600 measuring machine
Built prototype of the fiber plane prototype to test machine (4 fibers on a square 10 cm on a side)
Reproducible to 0.7 microns over 10 cm distances
Reproducible to 2 micron accuracy over full range of 20 “
Found worn guide bushings, replaced, trying again…

## Slide 28: Prototype of the Prototype Fiber Plane

## Slide 29: Precision Measuring Machine

## Slide 30: Measuring Machine Operation
Measuring machine runs under computer control, image displayed on computer screen
Magnification adjusted to display 120 micron diameter fiber to a 1 inch diameter on screen
Computer is programmed to measure 40 points on circumference of fiber
Computer fits circle to the 40 points, finds center

## Slide 31: Test measurements
Mounted fiber plate with 64 fibers on measuring machine
Repeated measurements of the 64 fibers 13 times
Rotated fiber plate 180o
Repeated 13 measurements of each of the 64 fibers

## Slide 32: Test measurements
Data reduction
Carried out transformation of each set of measurements to the first measurement of that orientation, also of rotated measurements to unrotated
Transformed images to first image using 3rd order polynomial
  X = a1+a2x+a3y+a4xy+a5x2+a6y2+a7x3+a8y3+a9x2y+a10xy2
  Y = b1+b2x+….
Calculated average x and y positions of the 13 measurements of each fiber for the unrotated measurements, same for rotated measurements
Calculated RMS of the 64 fibers of each measurement relative to the average of the same orientation (A)
Calculated the RMS of the 64 fibers of the rotated measurements relative to the average for each fiber in the unrotated measurements (B)

## Slide 33: A. Measuring Machine Repeatability
microns
microns
Number of Measurements
σx
σy
RMS of the 64 fibers in each measurement
with respect to the average of the 13
measurements in the same orientation

## Slide 34: B. Measuring Machine Precision
microns
microns
Number of Measurements
RMS of the 64 fibers of each measurement
of the rotated plate with respect to the average
of the 13 measurements unrotated
σx
σy

## Slide 35: Camera 8.9 meters from Fiber Plane

## Slide 36: Fiber Plane 8.9 meters from Camera

## Slide 37: Sewer Pipe Light Baffle

## Slide 38: Sewer Pipe Light Baffle

## Slide 39: Fiber View Camera Test Exposures
Took lots of exposures with the prototype camera
  Spread over many days
  Exposure times varied from 1 to 4 seconds
  f stops varied from f/5.6 to f/16
  Camera temperature varied from 5o C to 25o C
  Manual (eyeball) focusing
Processed data with Sextractor version 2.5.0
Transformed images to desired coordinate system using 3rd order polynomial
  X = a1+a2x+a3y+a4xy+a5x2+a6y2+a7x3+a8y3+a9x2y+a10xy2
  Y = b1+b2x+….

## Slide 40: Image of the Fiber Plane

## Slide 41: Blowup of an image

## Slide 42: Blowup of an image

## Slide 43: Blowup of an image

## Slide 44: Showing intensities

## Slide 45: FVC Gain Measurement
Took dark frames with FVC camera
  5 exposures each at three exposure times (0.1, 0.2, 0.5 seconds)
  About 2000 ADUs/pixel, independent of exposure time
  Averaged the 5 exposures at each exp time
Took flats(sheet of white paper in front of camera, out of focus) at each exposure time. Subtracted ave darks from flats pixel by pixel for each exp time.
CCD read out via 2 channels, right and left. Selected two postage stamps 100x100 pixels each, one on right, one left
Calculated ave ADU/pixel and RMS around ave for 10,000 pixels
Plotted (RMS)2 vs ave ADU/pixel, slope is gain

## Slide 46: Right side, 0.5 sec exp, ADU/pixel
Ave = 21,595 ADU/pixel
σ = 220.4
Sqrt(21,595)=147

## Slide 47: FVC Gain Measurement
Average ADU/pixel
(RMS)2
~ 2 ADUs/electron
Taking the pixel by pixel sensitivity variation
of (0.6±0.1) % into account

## Slide 48: Centroid Precision on CCD
Number of Exposures
millipixels
millipixels
20 different exposures, camera dithered by a few pixels only
Each entry is the RMS of 64 fiber position measurements
σx
σy

## Slide 49: Status last progress report Problem we are working on
millipixels
millipixels
Number of Exposures
σx
σy
Compare two sets of images : 10 measurements
with the camera in one fixed position, compared
to ten images with the camera moved laterally

## Slide 50: Status last progress report FVC Camera to OGP Meas Mach Comparison

## Slide 51: Detailed Distortion Correction
3rd order polynomial transformation not sufficient to take care of system distortions (lens distortion, CCD registration)
One option would be to go to higher order transformations
Distortion pattern looks irregular, so instead chose to create a “distortion map” (mask)
Took 314 exposures of the 64 fibers, moving image on CCD around to cover a 4000 x 4000 pixel area of the CCD 314x64 ~ 20,000 fiber images
Transformed all of the fiber measurements onto the OGP measuring machine coords, calculated the deviations Δx and Δy for each fiber measurement from their OGP measured positions.

## Slide 52: Radial lens distortion
0 15 30 45 60
centimeters on fiber plane

## Slide 53: Detailed Distortion Correction
Divided the area of the CCD into 10 pixel x 10 pixel area elements
In each area element calculated the average Δx and the average Δy displacements of the fiber positions from the OGP measurements within a 50 pixel radius
These ave Δx and Δy displacements in each area element on the CCD constitute the distortion map (calculated this once, stored for future use)

## Slide 54: Distortion Map Quality Check
Took 16 new exposures (4 sec exposures at f/16, ~30,000 e/fiber image) which were not used in creating the distortion map, varying positions on the CCD by 50 to 100 pixels
Found centroids of fiber images, corrected each centroid using the distortion map
Used 3rd order polynomial transformation to transform each exposure onto the OGP meas machine coords, calculated deviations from OGP measured positions
Also compared the 16 exposures to each other
  Called first exposure the reference image
  Transformed the other 15 onto the reference image
  Calculated RMS in x and y for the 64 fibers in each exposure, divided by √2

## Slide 55: FVC Image to Image Comparison
RMS in x and y for 64 fibers in each of 15 FVC exposures
milipixels
No of exposures

## Slide 56: Effect of correction mask
FVC Images compared to OGP meas machine
Δx, Δy ~ 2.5 microns on Fiber Plane!!

## Slide 57: Uncorrected two dimensional deviation FVC – Meas Machine

## Slide 58: Corrected two dimensional deviation (FVC – Meas Machine)

## Slide 59: Fiber position error on fiber plane
Two dimensional Δr ( FVC – Measuring Machine pos)

## Slide 60: A few hardware improvements

## Slide 61: Hardware improvements since last progress report
Implemented new light source using a 16” integrating sphere with 6 monochromatic LED’s, 521±25 nm, 100mW each
2 second exposures at f/16 give about 30,000 electrons per fiber image on CCD
Fiber plane
LED’s

## Slide 62: Alternate way of making a Distortion Map
A method independent of the fake fiber plane and the OGP measuring machine
Prof Van Altena’s group at Yale has a measuring machine (the PDS) used to measure star locations on photographic plates to ~ 1 micron precision
Have many 17” x 17” plates with ~4000 star images each
Measured star images on one plate on the DDS meas mach
Photographed the same plate 16 times with the FVC in different locations to cover the full 6000x6000 pixel area of the CCD
Used this data to construct a new distortion map
Used this distortion map to transform FVC images of the fake fiber plane and compared to the OGP measured locations

## Slide 63: The PDS Measuring Machine
Photometric Data System (Perkin-Elmer)

## Slide 64: Distortion maps
Original OGP
Full frame OGP
Full frame PDS

## Slide 65: Comparison of distortion map results
f/16 FVC data
f/5.6 and 7.3 FVC data
dy (μm)
dy (μm)

## Slide 66: Comments on the Observed Distortions
Distortion maps indicate distortion RMS of 40 milipixels, or 1/25 of a pixel, which is 0.24 microns on CCD, with a ~500 pixel periodicity
Pattern not suggestive of lens distortion, more likely to be distortions in the registration on the CCD
Talked to Kodak (now Truesense Imaging).
  Their engineer thought that the nominal 6 micron pixel size could vary on the CCD between 5.99985 and 6.00015 microns
  If the pixels were on the small side for 500 consecutive pixels and on the big side for 500 pixel this could make a variation of 0.00015x500=0.075 or ± 1/13 of a pixel.
  Is this a plausible explanation?

## Slide 67: A minor accident with an interesting result
Developed a crack in the camera window, CCD collected frost
Returned camera to FLI, replaced window
Camera works fine now
Returned to make distortion maps

## Slide 68: Distortion maps with new window
f/5.6
f/10.4
f/16
Distortion maps with star plates and PDS meas machine

## Slide 69: Distortion Map Comparison
Old Window
New Window
Distortion mostly due to the window!!!

## Slide 70
Distortions by the Window
Window flatness specified as 5 λ per inch
window
CCD
n
Θ2 = (n-1) Θ1 ~½ Θ1
Θ1 ~ 1μ/2500μ=0.4 mrad
Θ2 ~ 0.2 mrad
Δx = 0.2 mradx9mm~2μ
~0.3 pixels

## Slide 71: Results with the new window
FVC data at f/16, PDS maps
σ = 2.5μ
σ = 2.5μ

## Slide 72: Resolution on CCD
21 FVC exposures of the fiber plane with 64 fibers each
Exposures moved around to cover full frame of CCD
Each entry is the RMS of 64 fibers on one exposure
millipixels
No.of exposures
33 mpix goal
For all 21x64=
1344 meas
X rms = 15.7 mpix
Y rms = 16.5 mpix
f/16

## Slide 73: Resolution on Fiber Plane
No. of exposures
Microns on fiber plane
21 FVC exposures of the fiber plane with 64 fibers each
Exposures moved around to cover full frame of CCD
Each entry is the RMS of 64 fibers on one exposure
5 μ goal
For all 21x64=
1344 meas
X rms = 2.16 μ
Y rms = 2.28 μ
f/16

## Slide 74: 2 dimensional resolution on fiber plane
Delta x (microns)
Delta y (microns)
21x64 = 1344 measmnts
92% are within 5 μ circle
f/16

## Slide 75: Getting Consistent Results
Distortion Map | Old Window | New Window
Fibers w/ OGP | σ ~ 2.5μ | 
Stars w/ PDS | σ ~ 2.5μ | σ ~ 2.5μ
Took lots more data
Covered full 6000 x 6000 area of the CCD
Varied systematics
Distortion map with a third order polynomial
transformation able to correct for aberrations

## Slide 76: Exploring the Parameter Space
It looks like the prototype fiber view camera is able to provide better then the required 5μ resolution on the fiber plane. We next turn to explore the dependence on various parameters:
The FWHM of the fiber image on the CCD
The light level coming out of the fibers
Sanity checks on the lens: depth of field, diffraction limit
The f stop of the camera lens
Effects of turbulence due to thermal gradients
Effect of off-angle illumination of the fibers

## Slide 77: Dependence on FWHM of image on CCD
Image FWHM on CCD (pixels)
Resolution on Fiber Plane (μ)

## Slide 78: Dependence on the Flux
Flux in image (units of 103 ADU counts)
Resolution on fiber plane (microns)
2 ADU counts = 1 electron

## Slide 79: Dependence on the Flux
Flux in image (units of 103 ADU counts)
Resolution on fiber plane (microns)
2 ADU counts = 1 electron
FWHM/√N

## Slide 80: FVC Gain Measurement
Average ADU/pixel
(RMS)2
~ 2 ADUs/electron
Taking the pixel by pixel sensitivity variation
of (0.6±0.1) % into account

## Slide 81: Dependence on the Flux
There is a wide plateau between 20,000 and 80,000 ADU counts that give resolution better the 2.5μ at f/16.
With 2 ADU counts/electron this means between 10,000 and 40,000 e per image on CCD, so 25,000 is a good guess of where we should be.
Our original estimate was 75,000 e per image at f/5.6. With the smaller aperture at f/16 will still need more light out of the fibers.

## Slide 82: Depth of Field at Various f stops
Focus (object distance in cm)
FWHM on CCD (pixels)

## Slide 83: Diffraction Limit vs. fstop
Parameter | f/5.6 | f/10.4 | f/16
Aperture dia d | 71 mm | 39 mm | 25 mm
δΘ = λ/d | 8.5x10-6 rad | 15.4 x 10-6 rad | 24 x 10-6 rad
f x δΘ | 3.4 μ | 6.16 μ | 9.6 μ
Δr (pixels) | 0.6 | 1.0 | 1.6
Minimum image size due to diameter of the fibers:
120μ demagified by 25 gives 4.8μ image dia
or 0.8 pixels on CCD. We realized this early on
and said that we would have to defocus the image.
The diffraction limit at f/16 is doing this for us over
a wide range of focus positions,making f/16 very
robust. At lower f stops things get more delicate…

## Slide 84: Dependence on Lens f Stop
FVC Exposure | No Mask | f/5.6 Mask | f/10.4 Mask | f/16 Mask
f/5.6 | X rms 5.03 Y rms 4.19 | 5.29 4.62 | 4.92 3.94 | 4.87 4.01
f/10.4 | X rms 5.46 Y rms 4.52 | 5.56 5.27 | 5.16 4.59 | 5.19 4.53
f/16 | X rms 2.88 Y rms 2.55 | 3.11 3.11 | 2.36 2.25 | 2.16 2.28
Poor performance at f/5.6 and f/10.4 due to
some slight misalignement someplace
as we will see later
x and y resolutions on fiber plane in microns

## Slide 85: Peak Flux vs FWHM
FWHM on CCD (pixels )
Peak Flux

## Slide 86: Spread of FWHM on one Frame at f/10.4
y
Clearly some systematic problem
Could be misalignment in the
system someplace
Will look into this further
Not an issue at f/16 since there
we are diffraction limited and have
a very broad depth of field

## Slide 87: Depth of Field at Various f stops
Focus (object distance in cm)
FWHM on CCD (pixels)

## Slide 88: Possible Effect of Turbulence
Did one test to simulate turbulence in the Mayall telescope
Opened shades on windows, let the sunshine in. This heated up the sunnyside of the sewer pipe by up to 5oC
Took 20 exposures at each f stop at ΔT = 1oC and at ΔT = 5oC
No noticeable increase in the image FWHM at f/16
f stop | Previous resolution | ΔT = 1oC | ΔT = 5oC
f/16 | 2.22 μ | 2.06 μ | 2.49 μ
f/10.4 | 4.88 μ | 5.45 μ | 4.94 μ
f/5.6 | 4.95 μ | 5.85 μ | 4.65 μ

## Slide 89: Sewer Pipe Light Baffle

## Slide 90: Resolution VS FWHM at f/5.6

## Slide 91: Resolution vs FWHM at f/10.4

## Slide 92: Resolution vs FWHM at f/16

## Slide 93: Fiber Illumination Input Angle
Normal incidence
Sideways

## Slide 94: Hardware improvements since last progress report
Implemented new light source using a 16” integrating sphere with 6 monochromatic LED’s, 521±25 nm, 100mW each
2 second exposures at f/16 give about 30,000 electrons per fiber image on CCD
Fiber plane
LED’s

## Slide 95: Fiber Illumination Input Angle
No effect on resolution on fiber plane
Do all fibers move sideways coherently? Can not tell from present setup—integrating sphere and fiber plane on the same table (not particularly sturdy). The fiber plane as a whole seems to move around randomly between exposures by as much as 50 to 100 microns
Illumination Angle | σ (x) microns | σ (y) microns
Sideways (4 exposures) | 2.09 | 2.35
Normal (6 exposures) | 2.11 | 2.47
Sideways (4 exposures) | 2.15 | 2.08

## Slide 96: Fiber Plane Prototype 64 fibers total

## Slide 97: The Next Few Steps…
Double check, triple check results obtained so far by taking more images under a variety of conditions. Done
Is there a better way to create the distortion map? Done
Hardware improvements
  Install λ /15 optical flat window
  Present readout time is about 5 seconds, limited by USB cable between camera and computer.
  CCD capable of 1 sec readout with ethernet cable
Continue software development
  Optimize transformation equations Done
  Build in lens and CCD distortion maps Done
  Investigate behaviour at f/10.4 and f/5.6
  How many fiducials do we need
  Software running time to process an image (would like 1 second)

## Slide 98
R&D WBS 2.3 Fiber View Camera R&D Milestones
Feb 28, 2012 Fiber View Camera Pre-Conceptual Req Tech Note 1 Done Feb 28, 2012
Working Decision on FVCamera location, Tech Note 2 Done Feb 28, 2012
March 31, 2012 CCD Review Tech Note 4 Done March 27, 2012
Lens Review Tech Note 4 Done March 27, 2012
Pre-conceptual FVCamera Design Tech Note 3 Done March 19, 2012
April 30, 2012 Choice of CCD, lens, and electronics for Prototype Camera
Tech Note 5 Done March 20, 2012
May 31, 2012 Place Orders for CCD, Lens, Electronics for Prototype Done March 31, 2012
Aug 30,2012 Complete assembly of Prototype Camera Done May 1, 2012
Sept 30,2012 Centroid Test Setup Design Done April 15, 2012
Dec 31, 2012 Report on Prototype Camera Testing Tech Note 6 Done June 12, 2012
Jan 31, 2013 Complete test setup, start centroid tests Done June7, 2012
Feb 28, 2013 FVC Conceptual Requirements Document , Tech Note 7 Done Feb 24, 2013
March 31, 2013 Detector and Optics Selection for Conceptual Design Camera Done March 26, 2013
Tech Note 8
April 30, 2013 Conceptual Fiber View Camera Design Document Done April 25, 2013
May 31, 2013 Complete centroid tests Done May 30, 2013
Aug 30, 2013 Centroid Test Result